Veřejná zakázka s názvem:

Vertical electron beam evaporator with accessories for MBE system (OP JAK - FerrMion)

The subject of the procurement process is the supply of 3 electron guns (e-beam evaporators) with accessories (flux controller, shutter, power supply) for the MBE system* (Molecular Beam Epitaxy).

3 electron guns with accessories are required to vaporize the higher melting point metals required for the synthesis of hybrid ferroic materials. The electron guns will be mounted on the bottom of the MBE vacuum chamber and will point towards the substrate in the center of the chamber.

The growth of the hybrid films should take place at UHV pressure in order to guarantee the highest possible deposition purity and thereby limit contamination. The power of the electron guns should be adjustable so that it can also have a low deposition rate to control and adjust the deposition parameters during the first growth steps.

* The MBE system will be used for the preparation of hybrid nanostructures based on fullerenes and ferroic nanoparticles. The MBE system will consist of a large-volume vacuum chamber with a connected system of active elements, a sample manipulator, a thickness monitor and a system for linear sample transfer.

The contracting authority sets the following minimum technical conditions:

| Subject | Parameter / Required technical and functional properties, value | The specific parameters of the offered solution or the supplier indicates whether it meets YES / NO | Offered solution - exact type, part number, leading to identification (or website link) | Unit offer price in EUR without VAT | Total offer price in EUR without VAT |
| --- | --- | --- | --- | --- | --- |
| Electron guns (3 pcs) | UHV compatible |  |  | ……..EUR | ……..EUR |
|  | Mounting flange DN35CF-DN40CF (2.75" OD) or DN63CF (4.50” OD) |  |  |  |  |
|  | Suitable for evaporation of refractory material (W, Ta, Mo), C, Pt, Cr, Pt, Fe, etc. from rod (diameter from 1 up to 6 mm), and Au, Ag, Ni, etc. from crucible (with volume between 0.25 cc and 1 cc) |  |  |  |  |
|  | Flux control system integrated with the e-gun |  |  |  |  |
|  | E-beam power 50W-600W required range, or more |  |  |  |  |
|  | Internal water-cooling |  |  |  |  |
|  | Manual shutter (with a rotary feedthrough) |  |  |  |  |
|  | Compatible power supply for 19” rack with cables |  |  |  |  |
|  | Warranty period min. 12 months |  |  |  |  |
| Total offer price (3 pieces) including transport (in EUR without VAT) |  |  |  |  | ……..EUR |
| Estimated value of the subject of performance including transport to the headquarter (in CZK without VAT) |  |  |  |  | 1 115 000 CZK |

Note: the supplier must fill in all fields highlighted in yellow

___________________________________________

signature of an authorized person on behalf of the supplier